## Slide 1
BROOKS mass flow control
Number | Color | Description | Arduino
1 | Red-blue | Setpoint common | GND
2 | White-green | Flow Signal 0-5 V | A0
5 | Blue | Power Supply +13.5 to +27 VDC | -
8 | Green | Setpoint Signal 0 - 5 V | Pin 9
9 | Gray-brown | Power Supply common | GND
10 | Brown-green | Flow Signal output common | GND
BROOKS
Arduino
Custom cable
Twofast-rpi3-4
USB Serial
Database (twofast-rpi3-0)
LAN cable to network
Table: flow_meter_control
Columns: (setpoint_voltage, id)
Control is performed via the dash control, started via LTH_Neutimag\hkromer\17_github\dash_NG\NG_control_v2 and there the index.py with python 3.7.3. It is working in the conda environment “dash_readout”, i.e. on the MPC2278